## Slide 1
REAL-TIME PATIENT
MONITORING SYSTEM
Development on Raspberry Pi 5
Prepared by: Masrur Jamil Prochchhod Institution: Hochschule Hamm-Lippstadt (HSHL) Platform: Raspberry Pi 5 | Python 3 | Blynk IoT
HSHL Internship Project | 1

## Slide 2
PROJECT OVERVIEW
Goal
Build a Raspberry Pi 5 based patient monitoring prototype with local hardware output and cloud sync.
Sensors (Simulated)
Heart Rate
Body Temperature
Physical Acceleration
All generated via sim.py — no physical sensors wired in this phase.
Real Hardware
HD44780 16×2 LCD
(GPIO, 4-bit mode)
Active Piezo Buzzer
(GPIO Pin 17)
Real-Time Patient Monitoring System | HSHL Internship | Masrur

## Slide 3
HARDWARE SETUP — Raspberry Pi 5 + GPIO
HD44780 16×2 LCD — 4-Bit Parallel Mode
RS
GPIO 22
E (Enable)
GPIO 18
D4
GPIO 25
D5
GPIO 24
D6
GPIO 23
D7
GPIO 12
Active Piezo Buzzer
GPIO Pin 17 | Direct Digital Drive (HIGH/LOW)
No PWM required — triggers instantly on threshold breach
Simulated Vital Parameters
Heart Rate
Range: 40 – 150 BPM
Tachycardia / cardiac monitoring
Body Temperature
Range: 35°C – 42°C
Fever / hypothermia detection
Acceleration (3-axis)
Range: 0 – 5 G
Fall detection / motion monitoring
Real-Time Patient Monitoring System | HSHL Internship | Masrur

## Slide 4
THE HARDWARE STRUGGLE — SD Card Boot Failure
⚠ The Pink Screen
Every boot attempt produced: → Solid pink / rainbow screen → Error messages then halt → Occasional partial boot → freeze → Rare successful boots that wouldn't repeat Days passed with no usable system.
Possible Causes Investigated:
Corrupted SD filesystem
Bootloader cannot read boot partition
Degraded SD sectors
Physical wear causing intermittent failures
Incompatible card speed
Card too slow for Pi 5's I/O demands
Underpowered supply
Pi 5 draws more current than earlier models
Bad OS image write
Initial flash incomplete or corrupted
✓ Resolution
After exhausting all other options, the SD card was replaced with a new high-quality Class 10 / U3 card.
The Pi 5 booted cleanly and consistently on every subsequent power cycle — immediately and definitively.
Key lesson: SD cards are a common point of failure on Pi 5. Use a Class 10/U3 card or NVMe SSD from the outset.
Real-Time Patient Monitoring System | HSHL Internship | Masrur

## Slide 5
SOFTWARE CHALLENGE A — Wayland Display Driver Conflict
Symptom
matplotlib plot window opened completely blank — or threw a display server exception immediately on launch. No plot content rendered at all.
Root Cause
RPi 5 migrated from X11 → Wayland desktop protocol. matplotlib defaulted to the QtAgg backend, which tried to open via X11. On a Wayland system, QT_QPA_PLATFORM mismatch caused the rendering pipeline to fail entirely.
Resolution
Force-switch matplotlib to the TkAgg backend before any other import. Tkinter canvas renders correctly under Wayland without the Qt/X11 dependency.
The Fix — one line before all other matplotlib code:
import matplotlib matplotlib.use('TkAgg') # Must be called BEFORE importing matplotlib.pyplot import matplotlib.pyplot as plt
Real-Time Patient Monitoring System | HSHL Internship | Masrur

## Slide 6
SOFTWARE CHALLENGE B — CSV Race Conditions & Memory Leaks
Problem 1: File Zeroing Race
sim.py used write mode ('w') — erasing the entire CSV on every loop. Reading scripts caught the file at 0 bytes, causing flat lines, index errors, and crashes.
Problem 2: Memory Leak
FuncAnimation appended data arrays on every frame with no size limit. Arrays grew indefinitely, consuming RAM until the system slowed and failed.
✓ Resolution: Append-and-Trim Pattern
# WRITE: append only — never overwrites the file with open(FILENAME, 'a', newline='') as f: writer.writerow([hr, tmp, acc, t_now, d_now]) # READ: moving window — discards corrupted frames try:
lines = list(csv.reader(f))
if len(lines) < 2: return
recent = lines[
-20:] # last 20 rows only except (ValueError, IndexError, csv.Error): pass
Real-Time Patient Monitoring System | HSHL Internship | Masrur

## Slide 7
FINAL SYSTEM ARCHITECTURE — Data Flow
sim.py (Simulator)
Generates HR, Temp, Accel
Append-and-Trim writes
sensor_data.csv
Shared broker — 50-row max
medical_system.py
16×2 LCD • Buzzer alarm
Blynk IoT mobile sync
live_plot.py
3 stacked subplots • Fixed axes
Background flash on threshold breach
☁ Blynk IoT Cloud — Mobile Dashboard
Real-Time Patient Monitoring System | HSHL Internship | Masrur

## Slide 8
LIVE VISUALIZATION — Threshold Monitoring
Heart Rate
Axis Range:
40 – 150 BPM
Alarm Threshold:
100 BPM
Alert:
Background → red
ax1.set_facecolor('#ffcccc' if hrs[-1] > 100 else 'white')
Body Temperature
Axis Range:
35°C – 42°C
Alarm Threshold:
38.5°C
Alert:
Background → red
ax2.set_facecolor('#ffcccc' if tmps[-1] > 38.5 else 'white')
Acceleration
Axis Range:
0 – 5 G
Alarm Threshold:
2.0 G
Alert:
Background → red
ax3.set_facecolor('#ffcccc' if accs[-1] > 2.0 else 'white')
Design principle: Fixed axis ranges (not auto-scale) prevent clinical variations from being visually compressed and missed. Each subplot also shows a dashed horizontal threshold line as a persistent reference even when values are within normal range.
Real-Time Patient Monitoring System | HSHL Internship | Masrur

## Slide 9
CONCLUSION & NEXT STEPS
Challenges Resolved:
✓
[Hardware] SD card pink screen — replaced card
✓
[Software] Wayland/Qt blank plot — TkAgg fix
✓
[Software] CSV race condition — append-and-trim
✓
[Software] Memory leak — moving-window reads
✓
[Software] Flat line rendering — defensive parsing
Next Steps:
→ Wire real sensors: MAX30102, DS18B20, MPU6050
→ Migrate to NVMe SSD for reliable storage
→ Replace CSV broker with MQTT messaging
→ Add multi-patient support & identity layer
→ Deploy on edge device with UPS backup
HSHL Internship | Masrur | Raspberry Pi 5 Patient Monitoring System