MEDICAL EQUIPMENT

Standards & Interoperability Study

HORIBA ABX Micros 60  ·  Abaxis Piccolo Xpress

Analysis of FHIR, HL7, LOINC Support & Generated Data

Hochschule Hamm-Lippstadt (HSHL)

Student: Masrur Jamil Prochchhod

Date:04 June 2026

# 1. Executive Summary

This study evaluates two widely used clinical laboratory analyzers — the HORIBA ABX Micros 60 hematology analyzer and the Abaxis Piccolo Xpress point-of-care chemistry analyzer — from the perspective of health data interoperability standards. Specifically, it examines their support for HL7 messaging, FHIR R4 compatibility, and LOINC coding of generated laboratory data.

Both instruments are designed for laboratory and point-of-care settings and generate structured clinical data that must be transmitted to Laboratory Information Systems (LIS) and Electronic Medical Record (EMR) systems. The degree to which they support modern health data standards directly impacts their integration with hospital IT infrastructure and their suitability as reference points for the Smartband V001 project.

| Criterion | ABX Micros 60 | Piccolo Xpress |
| --- | --- | --- |
| Device Type | Hematology Analyzer (CBC) | Point-of-Care Chemistry Analyzer |
| Manufacturer | HORIBA ABX SAS (HORIBA Medical) | Abaxis Inc. (now Abbott Point of Care) |
| HL7 Support | Yes — via LiteDM software (HL7 bridge) | Yes — native HL7 output |
| FHIR R4 Support | Indirect — via middleware/integration | Indirect — via middleware/integration |
| LOINC Coding | Supported via LIS integration | Supported via LIS/EMR integration |
| LIS Connectivity | RS-232 / Ethernet via LiteDM | USB / Ethernet bidirectional |
| Data Format | ASTM, HL7 (via LiteDM) | ASCII, ASTM, HL7, XML |
| Direct FHIR API | Not natively supported | Not natively supported |

2. HORIBA ABX Micros 60 — Hematology Analyzer

## 2.1 Product Overview

The HORIBA ABX Micros 60 is a compact benchtop hematology analyzer designed for small to medium clinical laboratories. It performs a Complete Blood Count (CBC) with 3-part leukocyte differential using electrical impedance technology and photometric hemoglobin measurement. It is part of HORIBA's ABX Micros Series, manufactured by HORIBA ABX SAS in France.

The analyzer processes up to 60 samples per hour in open tube mode and 55 samples per hour in closed tube mode. It requires only 10 µL of whole blood per cycle, making it suitable for pediatric and microsampling applications. The device uses stepper motor technology instead of a compressor, resulting in reduced noise and maintenance requirements.

## 2.2 Technical Specifications

| Parameter | Value |
| --- | --- |
| Throughput | 60 samples/hour (open tube) / 55 samples/hour (closed tube) |
| Sample Volume | 10 µL whole blood (CBC and CBC 3-part DIFF) |
| Measurement Method | Impedance (WBC, RBC, PLT) + Photometry (HGB at 555 nm) |
| HCT Measurement | Numeric integration |
| 3-Part DIFF Method | Impedance with specific lyse (LMG) |
| Parameters | 18 parameters: WBC, RBC, HGB, HCT, MCV, MCH, MCHC, PLT, MPV, PCT, PDW, RDW, LYM%, LYM#, MON%, MON#, GRA%, GRA# |
| Memory | Last result + histograms (expandable with optional smart cards: 78 results) |
| Data Output | RS-232 (unidirectional); LiteDM software for LIS/EMR connectivity |
| Display | 2-line LCD, 40 alphanumeric characters per line |
| Printer | Dot matrix printer |
| Power Supply | 90–240 VAC, 50–60 Hz, 230 VA maximum |
| Dimensions | 44 cm (H) × 36 cm (W) × 33 cm (D) |
| Weight | 14 kg |
| Operating Temp | 16–34°C, max 80% relative humidity |
| Certifications | UL 61010-1, CAN/CSA-C22.2 No. 61010-1 |

## 2.3 Generated Data — 18 CBC Parameters

The ABX Micros 60 generates the following 18 hematological parameters per blood sample:

| Parameter | Abbreviation | LOINC Code | Unit | Clinical Significance |
| --- | --- | --- | --- | --- |
| White Blood Cell Count | WBC | 6690-2 | 10³/µL | Infection, immune status, leukemia screening |
| Red Blood Cell Count | RBC | 789-8 | 10⁶/µL | Anemia, polycythemia screening |
| Hemoglobin | HGB | 718-7 | g/dL | Anemia, oxygen-carrying capacity |
| Hematocrit | HCT | 4544-3 | % | Anemia, dehydration assessment |
| Mean Corpuscular Volume | MCV | 787-2 | fL | Anemia classification (micro/macrocytic) |
| Mean Corp. Hemoglobin | MCH | 785-6 | pg | Anemia subtype classification |
| Mean Corp. Hgb Concentration | MCHC | 786-4 | g/dL | Hemoglobin concentration per cell |
| Red Cell Distribution Width | RDW | 788-0 | % | Anemia type differentiation |
| Platelet Count | PLT | 777-3 | 10³/µL | Clotting disorders, thrombocytopenia |
| Mean Platelet Volume | MPV | 32623-1 | fL | Platelet function, cardiovascular risk |
| Platelet Crit | PCT | 51637-7 | % | Platelet mass in blood |
| Platelet Distribution Width | PDW | 32207-3 | % | Platelet size variation |
| Lymphocyte % | LYM% | 736-9 | % | Viral infections, immune response |
| Lymphocyte # | LYM# | 731-0 | 10³/µL | Absolute lymphocyte count |
| Monocyte % | MON% | 5905-5 | % | Inflammatory conditions |
| Monocyte # | MON# | 742-7 | 10³/µL | Absolute monocyte count |
| Granulocyte % | GRA% | 764-1 | % | Bacterial infections |
| Granulocyte # | GRA# | 30180-4 | 10³/µL | Absolute granulocyte (neutrophil + eosinophil + basophil) |

## 2.4 HL7 Support

The ABX Micros 60 itself does not natively output HL7 messages. HL7 connectivity is provided through the optional LiteDM (Lite Data Manager) software, which acts as a middleware bridge between the analyzer and LIS/EMR systems.

The LiteDM v3 Data Manager software enables HL7-based connectivity to EMR systems. Confirmed compatible EMR interfaces include Athena, Flatiron, Epic, eClinicalWorks, and CureMD. The connection is established via an HL7 bridge that translates the analyzer's RS-232 ASTM output into HL7 v2.x messages for transmission to the LIS.

Key HL7 capabilities via LiteDM:

- Bidirectional HL7 v2.x messaging between analyzer and LIS/EMR

- OBR segment carries the test order information with CBC panel codes

- OBX segments carry individual parameter results with LOINC codes as observation identifiers

- ADT messages for patient demographic integration

- QC data transmission to the LIS for quality management

- Confirmed interfaces: Athena, Flatiron, Epic, eClinicalWorks, CureMD

## 2.5 FHIR Support

The ABX Micros 60 does not natively support FHIR R4. There is no built-in FHIR API or FHIR resource generation in the device or LiteDM software. FHIR integration requires an external middleware layer or integration engine.

FHIR integration pathway for ABX Micros 60:

- Step 1 — ABX Micros 60 sends ASTM/HL7 v2.x data to LiteDM via RS-232

- Step 2 — LiteDM forwards HL7 v2.x messages to LIS/EMR via HL7 bridge

- Step 3 — An integration engine (e.g., Mirth Connect, Azure Health Data) transforms HL7 v2.x to FHIR R4 DiagnosticReport and Observation resources

- Step 4 — FHIR R4 resources are stored in a FHIR server (e.g., HAPI FHIR) for downstream consumption

When transformed to FHIR R4, each CBC result would generate a DiagnosticReport resource containing multiple Observation resources — one per parameter — each coded with the appropriate LOINC code.

## 2.6 LOINC Support

LOINC coding is not performed by the ABX Micros 60 analyzer itself. LOINC codes are assigned during LIS/EMR integration, either manually configured in the LiteDM interface mapping or automatically applied by the receiving LIS. The table in Section 2.3 lists the correct LOINC codes for all 18 parameters generated by the analyzer.

In HL7 v2.x messages, the OBX-3 field (Observation Identifier) carries the LOINC code in the format: LOINC_code^parameter_name^LN. For example, the WBC result would appear as: 6690-2^Leukocytes^LN.

## 2.7 Connectivity Architecture

| Connection Type | Protocol | Direction | Notes |
| --- | --- | --- | --- |
| RS-232 Serial | ASTM E1394 | Unidirectional (out) | Base hardware output — sends patient and QC results to LIS |
| RS-232 via LiteDM | HL7 v2.x | Bidirectional | Via LiteDM bridge — full LIS/EMR integration |
| Ethernet via LiteDM | HL7 v2.x | Bidirectional | LAN connectivity for multi-analyzer configurations |
| Smart Card (opt) | Proprietary | Local storage | QC and patient data storage — 78 results capacity |
| LIS Integration | ASTM / HL7 v2.x | Bidirectional | Via LiteDM: Athena, Flatiron, Epic, eClinicalWorks, CureMD |
| FHIR R4 | Not native | — | Requires external HL7→FHIR transformation middleware |

# 3. Abaxis Piccolo Xpress — Point-of-Care Chemistry Analyzer

## 3.1 Product Overview

The Abaxis Piccolo Xpress is a fully automated, portable point-of-care blood chemistry analyzer designed for use in any patient-care setting. It provides lab-accurate clinical chemistry results within 12 minutes using only 100 µL (0.1 cc) of whole blood, serum, or plasma — equivalent to approximately 4 drops of blood.

The system uses patented single-use self-calibrating reagent discs, each containing all reagents required to perform a complete panel of up to 14 tests simultaneously. Abaxis was acquired by Zoetis in 2018; the human medicine product line (Piccolo Xpress) is now distributed in the USA by Abbott Point of Care. The device is widely used in primary care, urgent care, hospitals, military settings, and research facilities.

## 3.2 Technical Specifications

| Parameter | Value |
| --- | --- |
| Test Time | 12 minutes per complete panel |
| Sample Volume | 100 µL (0.1 cc) — whole blood, serum, or plasma |
| Sample Type | Lithium heparin whole blood, serum, or plasma |
| Available Panels | 17 pre-configured multi-analyte panels |
| Total Tests | 31 unique chemistry tests across all panels |
| CLIA Waived | 11 of 17 panels are CLIA waived |
| Measurement Method | Spectrophotometry — Xenon arc stroboscopic lamp, 9 wavelengths, 16-bit ADC |
| Reaction Temperature | 37°C ± 1°C (maintained by internal heater during analysis) |
| Quality Control | Built-in iQC® (intelligent Quality Control) system — automatic |
| Memory | Up to 5,000 patient and QC results |
| Display | VGA color touchscreen |
| Connectivity | 5 USB ports (bidirectional) + 1 Ethernet port |
| Data Formats | ASCII, ASTM, HL7, XML |
| Printer | Built-in thermal printer + external printer/LIS/EMR capable |
| Power Requirements | 100–240 VAC, 50–60 Hz or 15 VDC, 5.0A |
| Dimensions | 32.4 cm (H) × 14.9 cm (W) × 21.0 cm (D) |
| Weight | 5.1 kg (analyzer) + 0.7 kg (power adapter) |
| Operating Temp | 15–32°C, 8–80% humidity (non-condensing) |
| Operation Mode | Continuous, walk-away |

## 3.3 Generated Data — 31 Chemistry Tests

The Piccolo Xpress generates up to 31 distinct chemistry results across 17 panel configurations. The following table lists all 31 analytes with their LOINC codes and clinical significance:

| Analyte | Abbrev. | LOINC Code | Unit | Clinical Area |
| --- | --- | --- | --- | --- |
| Albumin | ALB | 1751-7 | g/dL | Liver function, nutrition |
| Alkaline Phosphatase | ALP | 6768-6 | U/L | Liver, bone disease |
| Alanine Aminotransferase | ALT | 1742-6 | U/L | Liver damage, hepatitis |
| Amylase | AMY | 1798-8 | U/L | Pancreatitis |
| Aspartate Aminotransferase | AST | 1920-8 | U/L | Liver, heart damage |
| Direct Bilirubin | DBIL | 1968-7 | mg/dL | Liver, bile duct disease |
| Total Bilirubin | TBIL | 1975-2 | mg/dL | Liver function, jaundice |
| Blood Urea Nitrogen | BUN | 3094-0 | mg/dL | Kidney function |
| Total Calcium | Ca | 17861-6 | mg/dL | Bone, parathyroid, kidney |
| Creatine Kinase | CK | 2157-6 | U/L | Heart attack, muscle damage |
| Chloride | Cl⁻ | 2075-0 | mmol/L | Electrolyte balance |
| Creatinine | CRE | 2160-0 | mg/dL | Kidney function |
| C-Reactive Protein | CRP | 1988-5 | mg/L | Inflammation, infection |
| Gamma-Glutamyl Transferase | GGT | 2324-2 | U/L | Liver function, alcohol use |
| Glucose | GLU | 2345-7 | mg/dL | Diabetes, metabolic disorders |
| Potassium | K⁺ | 2823-3 | mmol/L | Cardiac, renal, metabolic |
| Lactate | LAC | 2524-7 | mmol/L | Sepsis, tissue hypoxia |
| Lactate Dehydrogenase | LDH | 2532-0 | U/L | Tissue damage, hemolysis |
| Magnesium | Mg | 2601-3 | mg/dL | Neuromuscular function |
| Sodium | Na⁺ | 2951-2 | mmol/L | Fluid balance, hyponatremia |
| Phosphorus | PHOS | 2777-1 | mg/dL | Bone, kidney, parathyroid |
| Total Carbon Dioxide | tCO₂ | 2028-9 | mmol/L | Acid-base balance |
| Total Protein | TP | 2885-2 | g/dL | Nutrition, liver, kidney |
| Uric Acid | UA | 3084-1 | mg/dL | Gout, kidney stones |
| Total Cholesterol | CHOL | 2093-3 | mg/dL | Cardiovascular risk |
| HDL Cholesterol | HDL | 2085-9 | mg/dL | Cardiovascular protection |
| LDL Cholesterol* | LDL* | 13457-7 | mg/dL | Cardiovascular risk (calc.) |
| Non-HDL Cholesterol* | nHDL* | 70204-3 | mg/dL | Cardiovascular risk (calc.) |
| Triglycerides | TRIG | 2571-8 | mg/dL | Lipid metabolism, cardiac |
| VLDL Cholesterol* | VLDL* | 13458-5 | mg/dL | Cardiovascular risk (calc.) |
| Cholesterol/HDL Ratio* | CHOL/HDL* | 9830-1 | ratio | Cardiovascular risk ratio (calc.) |

* Calculated result, not directly measured.

## 3.4 Panel Configurations

The 31 analytes above are organized into 17 pre-configured panel configurations:

| Panel Name | Catalogue # | Tests Included | CLIA Waived? |
| --- | --- | --- | --- |
| General Chemistry 6 | 400-0006 | ALT, AST, BUN, CRE, GGT, GLU | Yes |
| General Chemistry 13 | 400-0029 | ALB, ALP, ALT, AMY, AST, BUN, Ca, CRE, GGT, GLU, TBIL, TP, UA | Yes |
| BioChemistry Panel Plus | 400-0035 | ALB, ALP, ALT, AMY, AST, BUN, Ca, CRE, CRP, GGT, GLU, TP, UA | No |
| AmLyte 13 | 400-0041 | ALB, ALT, AMY, AST, BUN, Ca, CK, CRE, CRP, GLU, K+, Na+, TBIL | No |
| Basic Metabolic Panel | 400-0024 | BUN, Ca, Cl-, CRE, GLU, K+, Na+, tCO2 | Yes |
| Basic Metabolic Panel Plus | 400-0031 | BUN, Ca, Cl-, CRE, GLU, K+, LDH, Mg, Na+, tCO2 | Yes |
| Comprehensive Metabolic Panel | 400-0028 | ALB, ALP, ALT, AST, BUN, Ca, Cl-, CRE, GLU, K+, Na+, TBIL, tCO2, TP | Yes |
| Electrolyte Panel | 400-0022 | Cl-, K+, Na+, tCO2 | Yes |
| MetLyte 8 Panel | 400-0023 | BUN, CK, Cl-, CRE, GLU, K+, Na+, tCO2 | Yes |
| MetLyte Plus CRP | 400-0034 | BUN, CK, Cl-, CRE, CRP, GLU, K+, Na+, tCO2 | No |
| MetLac 12 Panel | 400-0037 | ALB, BUN, Ca, Cl-, CRE, GLU, K+, LAC, Mg, Na+, PHOS, tCO2 | No |
| Lipid Panel | 400-0025 | CHOL, CHOL/HDL*, HDL, nHDL*, LDL*, TRIG, VLDL* | Yes |
| Lipid Panel Plus | 400-0030 | ALT, AST, GLU, CHOL, CHOL/HDL*, HDL, nHDL*, LDL*, TRIG, VLDL* | Yes |
| Hepatic Function Panel | 400-0026 | ALB, ALP, ALT, AST, DBIL, TBIL, TP | No |
| Liver Panel Plus | 400-0003 | ALB, ALP, ALT, AMY, AST, GGT, TBIL, TP | Yes |
| Renal Function Panel | 400-0027 | ALB, BUN, Ca, Cl-, CRE, GLU, K+, Na+, PHOS, tCO2 | Yes |
| Kidney Check | 400-0033 | BUN, CRE | Yes |

## 3.5 HL7 Support

The Piccolo Xpress has native HL7 support built into the analyzer hardware. It can output data directly in HL7 format via its USB and Ethernet ports without requiring additional software.

The Piccolo Xpress supports the following data transmission formats natively:

- ASCII — plain text format for simple connectivity

- ASTM E1394 — the legacy laboratory instrument standard

- HL7 — Health Level 7 v2.x messaging for LIS/EMR integration

- XML — extensible markup language for structured data exchange

Connectivity is bidirectional via 5 USB ports and 1 Ethernet port. The device can receive patient demographic orders from a LIS (downlink) and transmit results back (uplink). Data parsing and translation may be required depending on the receiving system's HL7 version and message format requirements.

The Relaymed POC connectivity platform provides automated HL7 integration for the Piccolo Xpress specifically designed for physician office laboratories — eliminating the need for a full LIS by providing direct EMR connectivity with LOINC-coded results.

## 3.6 FHIR Support

The Piccolo Xpress does not natively output FHIR R4 resources. However, its native HL7 and XML output capabilities make it significantly easier to implement FHIR integration compared to devices with only proprietary outputs.

FHIR integration pathway for Piccolo Xpress:

- Step 1 — Piccolo Xpress outputs results in HL7 v2.x or XML format via USB or Ethernet

- Step 2 — Results are received by a POC connectivity platform (e.g., Relaymed, Cerner CareAware, Epic Rover) or middleware

- Step 3 — The middleware transforms HL7 v2.x ORU^R01 messages into FHIR R4 DiagnosticReport and Observation resources

- Step 4 — FHIR resources are posted to the hospital FHIR server for EHR integration

When transformed to FHIR R4, each Piccolo panel result generates a DiagnosticReport resource (representing the panel) containing Observation resources for each analyte, each coded with its LOINC code, UCUM unit, reference range, and abnormal flag.

## 3.7 LOINC Support

The Piccolo Xpress does not internally assign LOINC codes to its output. LOINC code assignment is performed by the receiving LIS, EMR, or POC connectivity platform during result mapping. The Relaymed platform, for example, automatically maps all 31 Piccolo analytes to their LOINC codes when transmitting to an EMR.

All 31 analytes generated by the Piccolo Xpress have established LOINC codes (listed in Section 3.3). The Piccolo's XML output format facilitates accurate LOINC mapping because the analyte names are transmitted as structured text rather than proprietary codes.

## 3.8 Connectivity Architecture

| Connection Type | Protocol | Direction | Notes |
| --- | --- | --- | --- |
| USB Port (×5) | ASCII / ASTM / HL7 / XML | Bidirectional | Connect to printer, keyboard, barcode scanner, computer, or EMR/LIS |
| Ethernet Port (×1) | ASCII / ASTM / HL7 / XML | Bidirectional | Network LIS/EMR connectivity |
| Built-in Printer | Thermal print | Output only | Adhesive-backed paper — patient results with QC data |
| Relaymed Platform | HL7 → EMR | Bidirectional | POC connectivity software — LOINC coded results to EMR |
| LIS Direct | HL7 v2.x / ASTM | Bidirectional | Compatible with all major LIS platforms |
| FHIR R4 | Not native | — | Requires external HL7→FHIR transformation middleware |

# 4. Comparative Standards Analysis

## 4.1 HL7 Support Comparison

| Feature | ABX Micros 60 | Piccolo Xpress |
| --- | --- | --- |
| Native HL7 Output | No — requires LiteDM software | Yes — native HL7 v2.x output |
| HL7 Version | HL7 v2.x via LiteDM bridge | HL7 v2.x (ASCII, ASTM, HL7, XML) |
| Message Direction | Bidirectional via LiteDM | Bidirectional (USB + Ethernet) |
| Physical Interface | RS-232 serial | USB (×5) + Ethernet (×1) |
| EMR Compatibility | Athena, Flatiron, Epic, eClinicalWorks, CureMD | All HL7-compatible EMR/LIS systems |
| Setup Complexity | Requires LiteDM software installation + HL7 bridge config | Plug-and-play with HL7-capable systems |
| Additional Software | LiteDM v3 (optional, sold separately) | Relaymed (optional POC platform) |

## 4.2 FHIR R4 Support Comparison

| Feature | ABX Micros 60 | Piccolo Xpress |
| --- | --- | --- |
| Native FHIR R4 API | No | No |
| FHIR via Middleware | Yes — via HL7→FHIR transformation | Yes — via HL7/XML→FHIR transformation |
| FHIR Resource Type | DiagnosticReport + Observation (×18) | DiagnosticReport + Observation (×31 max) |
| Easiest FHIR Path | LiteDM → Mirth Connect → HAPI FHIR | USB/Ethernet → Relaymed → FHIR server |
| FHIR Observation Count | 18 per blood sample | Up to 31 per panel (panel-dependent) |
| SMART on FHIR | Not applicable at device level | Not applicable at device level |

## 4.3 LOINC Code Coverage

| Feature | ABX Micros 60 | Piccolo Xpress |
| --- | --- | --- |
| Native LOINC Assignment | No — assigned by LIS/EMR | No — assigned by LIS/EMR |
| Total Analytes | 18 CBC parameters | 31 chemistry analytes |
| LOINC Codes Available | All 18 have established LOINC codes | All 31 have established LOINC codes |
| OBX-3 LOINC in HL7 | Via LiteDM mapping configuration | Via LIS/Relaymed mapping |
| Code System | LOINC (LN) | LOINC (LN) |
| UCUM Units | All parameters use UCUM units | All analytes use UCUM units |

## 4.4 Overall Standards Readiness Score

The following assessment summarizes each device's readiness for modern health data interoperability standards:

| Standard | ABX Micros 60 Score | Piccolo Xpress Score | Notes |
| --- | --- | --- | --- |
| HL7 v2.x | ★★★☆☆ (3/5) | ★★★★★ (5/5) | Piccolo has native HL7; Micros 60 needs LiteDM |
| FHIR R4 | ★★☆☆☆ (2/5) | ★★★☆☆ (3/5) | Neither native; Piccolo easier via XML output |
| LOINC Coding | ★★★☆☆ (3/5) | ★★★★☆ (4/5) | Both support LOINC via LIS; Piccolo via XML is cleaner |
| LIS Integration | ★★★★☆ (4/5) | ★★★★★ (5/5) | Piccolo's multi-format + Ethernet is more flexible |
| Data Richness | ★★★☆☆ (3/5) | ★★★★★ (5/5) | Piccolo generates 31 analytes across organ systems |
| Ease of Setup | ★★★☆☆ (3/5) | ★★★★☆ (4/5) | Piccolo plug-and-play; Micros 60 needs LiteDM config |
| Overall | ★★★☆☆ (3/5) | ★★★★☆ (4/5) | Piccolo more interoperable out of the box |

# 5. Relevance to Smartband V001 and FHIR Integration

Both the ABX Micros 60 and the Piccolo Xpress generate laboratory data that is directly relevant to the health monitoring goals of the Smartband V001 project. Their data output patterns serve as reference models for how clinical laboratory analyzers structure and transmit measurement data.

## 5.1 LOINC Code Reference for Smartband V001

Several LOINC codes used by these analyzers are also applicable to Smartband V001 sensor measurements:

| Smartband V001 Measurement | Relevant Analyzer Data | LOINC Code | Connection |
| --- | --- | --- | --- |
| Heart Rate (MAX30102) | Not measured by these analyzers | 8867-4 | Shared LOINC coding approach |
| SpO2 (MAX30102) | Not measured by these analyzers | 2708-6 | Shared LOINC coding approach |
| Glucose (future sensor) | Piccolo Xpress: GLU | 2345-7 | Direct LOINC reuse from Piccolo |
| CRP (inflammation marker) | Piccolo Xpress: CRP | 1988-5 | Direct LOINC reuse from Piccolo |
| Hemoglobin (derived) | ABX Micros 60: HGB | 718-7 | Reference for wearable Hgb estimation |
| Electrolytes (future) | Piccolo Xpress: Na+, K+, Cl- | Multiple | Reference for future electrolyte sensor |

## 5.2 Key Learnings for FHIR Implementation

- Both analyzers use HL7 v2.x ORU^R01 messages to transmit laboratory results — the same message type that would carry Smartband V001 data if connected via a gateway.

- Neither analyzer generates FHIR R4 natively — confirming that middleware-based FHIR transformation (as implemented in the Smartband V001 Flask server) is the industry standard approach.

- LOINC codes are universally used for laboratory result coding but are assigned by the LIS/middleware layer, not by the instrument itself — consistent with Smartband V001's FHIR gateway approach.

- The DiagnosticReport + Observation FHIR resource pattern used by both analyzer integrations is the same pattern the Smartband V001 FHIR server implements for its sensor readings.

- The Piccolo Xpress's XML output demonstrates that structured data formats (not just ASTM) are increasingly common in POC devices — supporting the Smartband V001's JSON/FHIR approach.

# 6. Conclusion

The HORIBA ABX Micros 60 and Abaxis Piccolo Xpress represent two generations of clinical analyzer connectivity philosophy. The Micros 60, an older design, relies on RS-232 and requires middleware (LiteDM) for HL7 connectivity, reflecting the legacy approach to laboratory instrument integration. The Piccolo Xpress, a more modern point-of-care design, offers native HL7, ASTM, XML, and ASCII output via USB and Ethernet, reflecting the industry's shift toward more open, flexible connectivity.

Neither device natively supports FHIR R4 — which is consistent with the broader laboratory instrument market, where FHIR integration is almost universally implemented at the middleware/LIS layer rather than at the instrument itself. Both devices support LOINC coding through their LIS integration pathways, with all generated analytes having established LOINC codes.

For the Smartband V001 project, these analyzers serve as important reference points. They demonstrate that the middleware-based FHIR transformation approach implemented in the Smartband V001 Flask server is the correct and industry-standard method. The LOINC codes for analytes like glucose, CRP, and electrolytes — generated by the Piccolo Xpress — are directly reusable for future Smartband V001 sensor expansions.