## Slide 1
MEDICAL EQUIPMENT
Standards & Interoperability Study
HORIBA ABX Micros 60
Abaxis Piccolo Xpress
Analysis of FHIR R4, HL7, and LOINC Support
Student: Masrur Jamil Prochchhod Institution: Hochschule Hamm-Lippstadt (HSHL) Context: Reference study for Smartband V001 FHIR integration
Scope: HL7 messaging support · FHIR R4 compatibility · LOINC coding · LIS connectivity · Smartband V001 relevance
HSHL Internship Project · June 2026

## Slide 2
DEVICE OVERVIEW — At a Glance
HORIBA ABX Micros 60
Type
Hematology Analyzer (CBC)
Manufacturer
HORIBA ABX SAS (France)
Throughput
60 samples/hour (open tube)
Sample Vol.
10 µL whole blood
Method
Impedance + Photometry (HGB)
Parameters
18 CBC parameters
HL7
Via LiteDM middleware bridge
Connectivity
RS-232 serial output
FHIR
Indirect via middleware
Abaxis Piccolo Xpress
Type
Point-of-Care Chemistry Analyzer
Manufacturer
Abaxis (now Abbott Point of Care)
Test Time
12 minutes per panel
Sample Vol.
100 µL (whole blood/serum/plasma)
Method
Spectrophotometry (9 wavelengths)
Parameters
31 chemistry analytes / 17 panels
HL7
Native HL7 v2.x output
Connectivity
5× USB + 1× Ethernet
FHIR
Indirect via middleware
Medical Equipment Standards Study · ABX Micros 60 & Piccolo Xpress · HSHL 2026

## Slide 3
ABX MICROS 60 — 18 CBC Parameters with LOINC Codes
Abbrev.
LOINC
Unit
Parameter
Abbrev.
LOINC
Unit
Parameter
WBC
6690-2
10³/µL
White Blood Cell Count
RBC
789-8
10⁶/µL
Red Blood Cell Count
HGB
718-7
g/dL
Hemoglobin
HCT
4544-3
%
Hematocrit
MCV
787-2
fL
Mean Corpuscular Volume
MCH
785-6
pg
Mean Corp. Hemoglobin
MCHC
786-4
g/dL
Mean Corp. Hgb Concentration
RDW
788-0
%
Red Cell Distribution Width
PLT
777-3
10³/µL
Platelet Count
MPV
32623-1
fL
Mean Platelet Volume
PCT
51637-7
%
Platelet Crit
PDW
32207-3
%
Platelet Distribution Width
LYM%
736-9
%
Lymphocyte %
LYM#
731-0
10³/µL
Lymphocyte #
MON%
5905-5
%
Monocyte %
MON#
742-7
10³/µL
Monocyte #
GRA%
764-1
%
Granulocyte %
GRA#
30180-4
10³/µL
Granulocyte #
ABX Micros 60 · 18 CBC Parameters · All with established LOINC codes

## Slide 4
PICCOLO XPRESS — 31 Chemistry Analytes with LOINC Codes
Abbr.
LOINC
Unit
Analyte
ALB
1751-7
g/dL
Albumin
ALP
6768-6
U/L
Alkaline Phosphatase
ALT
1742-6
U/L
Alanine Aminotransferase
AMY
1798-8
U/L
Amylase
AST
1920-8
U/L
Aspartate Aminotransferase
DBIL
1968-7
mg/dL
Direct Bilirubin
TBIL
1975-2
mg/dL
Total Bilirubin
BUN
3094-0
mg/dL
Blood Urea Nitrogen
Ca
17861-6
mg/dL
Total Calcium
CK
2157-6
U/L
Creatine Kinase
Cl⁻
2075-0
mmol/L
Chloride
Abbr.
LOINC
Unit
Analyte
CRE
2160-0
mg/dL
Creatinine
CRP
1988-5
mg/L
C-Reactive Protein
GGT
2324-2
U/L
Gamma-Glutamyl Transferase
GLU
2345-7
mg/dL
Glucose
K⁺
2823-3
mmol/L
Potassium
LAC
2524-7
mmol/L
Lactate
LDH
2532-0
U/L
Lactate Dehydrogenase
Mg
2601-3
mg/dL
Magnesium
Na⁺
2951-2
mmol/L
Sodium
PHOS
2777-1
mg/dL
Phosphorus
Abbr.
LOINC
Unit
Analyte
tCO₂
2028-9
mmol/L
Total Carbon Dioxide
TP
2885-2
g/dL
Total Protein
UA
3084-1
mg/dL
Uric Acid
CHOL
2093-3
mg/dL
Total Cholesterol
HDL
2085-9
mg/dL
HDL Cholesterol
LDL*
13457-7
mg/dL
LDL Cholesterol (calc.)
nHDL*
70204-3
mg/dL
Non-HDL Cholesterol (calc.)
TRIG
2571-8
mg/dL
Triglycerides
VLDL*
13458-5
mg/dL
VLDL Cholesterol (calc.)
CHOL/HDL*
9830-1
ratio
Cholesterol/HDL Ratio (calc.)
Piccolo Xpress · 31 Chemistry Analytes · All with established LOINC codes · * = Calculated result

## Slide 5
HL7 & FHIR SUPPORT COMPARISON
HL7 Support:
Feature
ABX Micros 60
Piccolo Xpress
Native HL7 Output
No — requires LiteDM middleware
Yes — native HL7 v2.x output
HL7 Version
HL7 v2.x via LiteDM bridge
HL7 v2.x (ASCII, ASTM, HL7, XML)
Message Direction
Bidirectional via LiteDM
Bidirectional (USB + Ethernet)
Physical Interface
RS-232 serial
5× USB + 1× Ethernet
EMR Compatibility
Athena, Flatiron, Epic, eClinicalWorks, CureMD
All HL7-compatible EMR/LIS systems
Setup Complexity
Requires LiteDM install + HL7 bridge config
Plug-and-play with HL7-capable systems
FHIR R4: Neither device supports it natively. Both require middleware (Mirth Connect / Relaymed) to transform HL7 v2.x → FHIR R4 DiagnosticReport + Observation resources.
Medical Equipment Standards Study · ABX Micros 60 & Piccolo Xpress · HSHL 2026

## Slide 6
STANDARDS READINESS SCORES
HL7 v2.x
3/5
5/5
Piccolo has native HL7; Micros 60 needs LiteDM
FHIR R4
2/5
3/5
Neither native; Piccolo easier via XML output
LOINC Coding
3/5
4/5
Both support LOINC via LIS; Piccolo XML is cleaner
LIS Integration
4/5
5/5
Piccolo multi-format + Ethernet is more flexible
Data Richness
3/5
5/5
Piccolo 31 analytes vs Micros 60's 18 CBC params
Ease of Setup
3/5
4/5
Piccolo plug-and-play; Micros 60 needs LiteDM config
Overall
3/5
4/5
Piccolo more interoperable out of the box
ABX Micros 60
Piccolo Xpress
Medical Equipment Standards Study · ABX Micros 60 & Piccolo Xpress · HSHL 2026

## Slide 7
PICCOLO XPRESS — 17 Panel Configurations
Panel Name
Cat. #
Tests Included
CLIA
General Chemistry 6
400-0006
ALT, AST, BUN, CRE, GGT, GLU
General Chemistry 13
400-0029
ALB, ALP, ALT, AMY, AST, BUN, Ca, CRE, GGT, GLU, TBIL, TP, UA
Basic Metabolic Panel
400-0024
BUN, Ca, Cl⁻, CRE, GLU, K⁺, Na⁺, tCO₂
Comprehensive Metabolic
400-0028
ALB, ALP, ALT, AST, BUN, Ca, Cl⁻, CRE, GLU, K⁺, Na⁺, TBIL, tCO₂, TP
Electrolyte Panel
400-0022
Cl⁻, K⁺, Na⁺, tCO₂
Lipid Panel
400-0025
CHOL, CHOL/HDL*, HDL, nHDL*, LDL*, TRIG, VLDL*
Lipid Panel Plus
400-0030
ALT, AST, GLU + full lipid panel
Hepatic Function Panel
400-0026
ALB, ALP, ALT, AST, DBIL, TBIL, TP
Liver Panel Plus
400-0003
ALB, ALP, ALT, AMY, AST, GGT, TBIL, TP
Renal Function Panel
400-0027
ALB, BUN, Ca, Cl⁻, CRE, GLU, K⁺, Na⁺, PHOS, tCO₂
MetLyte 8
400-0023
BUN, CK, Cl⁻, CRE, GLU, K⁺, Na⁺, tCO₂
MetLyte Plus CRP
400-0034
BUN, CK, Cl⁻, CRE, CRP, GLU, K⁺, Na⁺, tCO₂
MetLac 12
400-0037
ALB, BUN, Ca, Cl⁻, CRE, GLU, K⁺, LAC, Mg, Na⁺, PHOS, tCO₂
BMP Plus
400-0031
BUN, Ca, Cl⁻, CRE, GLU, K⁺, LDH, Mg, Na⁺, tCO₂
BioChemistry Panel Plus
400-0035
ALB, ALP, ALT, AMY, AST, BUN, Ca, CRE, CRP, GGT, GLU, TP, UA
AmLyte 13
400-0041
ALB, ALT, AMY, AST, BUN, Ca, CK, CRE, CRP, GLU, K⁺, Na⁺, TBIL
Kidney Check
400-0033
BUN, CRE
CLIA Waived
Not waived
11 of 17 panels are CLIA waived

## Slide 8
RELEVANCE TO SMARTBAND V001 & FHIR INTEGRATION
LOINC Code Reference for Smartband V001:
Smartband Measurement
Analyzer Source
LOINC
Connection
Heart Rate (MAX30102)
Not in these analyzers
8867-4
Shared LOINC coding approach
SpO2 (MAX30102)
Not in these analyzers
2708-6
Shared LOINC coding approach
Glucose (future sensor)
Piccolo Xpress: GLU
2345-7
Direct LOINC reuse from Piccolo
CRP (inflammation)
Piccolo Xpress: CRP
1988-5
Direct LOINC reuse from Piccolo
Hemoglobin (derived)
ABX Micros 60: HGB
718-7
Reference for wearable Hgb estimation
Electrolytes (future)
Piccolo Xpress: Na⁺, K⁺, Cl⁻
Multiple
Reference for future electrolyte sensor
Key Learnings for Smartband V001 FHIR Implementation:
Both analyzers use HL7 v2.x ORU^R01 messages — same type that would carry Smartband V001 data via a gateway
Neither generates FHIR R4 natively — confirming middleware-based FHIR transformation is the industry standard
LOINC codes assigned by LIS/middleware layer, not by the device — consistent with Smartband V001's FHIR gateway approach
DiagnosticReport + Observation FHIR pattern (used by both) is the same pattern Smartband V001 FHIR server implements
Medical Equipment Standards Study · ABX Micros 60 & Piccolo Xpress · HSHL 2026

## Slide 9
CONCLUSION
ABX Micros 60 ★★★ (3/5)
Older RS-232 based design
HL7 requires LiteDM middleware
18 CBC parameters — all LOINC coded
FHIR via Mirth Connect middleware
Reliable but requires extra config
Piccolo Xpress ★★★★ (4/5)
Modern USB + Ethernet connectivity
Native HL7, ASTM, XML, ASCII output
31 analytes / 17 panels — all LOINC coded
FHIR via Relaymed or middleware
Plug-and-play — more interoperable
Neither device natively supports FHIR R4 — confirming that middleware-based FHIR transformation (as in Smartband V001's Flask server) is the correct industry-standard approach.
Medical Equipment Standards Study · Masrur Jamil Prochchhod · HSHL · June 2026